## The Changes here are not exactl
Requirement ID | Deleted Text | Added Text | Merged Changed Text
RLM.03955 | a process plan for implementing back-up capability... to purchase and deliver soon enough... to have an emergency procedure to provide backup units... At least annually to ensure their... using equipment is located at another laboratory | The laboratory has a process plan for utilizing back-up temperature-dependent equipment... a back-up process must be implemented as procedurally described to avoid loss of contents... The backup units must be tested at plan must address emergent situations (eg. disaster)... At least annually, the laboratory must assess backup equipment for functionality if neededand capacity and check the contact personnel list for accuracy... If the backup plan involves using equipment is located at another laboratory or transferring specimens toif storage occurs at another laboratory, there must be a written agreement between the laboratories | The requirement for temperature-dependent equipment back-up plans has been revised to focus on utilizing back-up equipment, addressing emergent situations like disasters, and assessing the functionality and capacity of the back-up equipment annually.
RLM.10166 | N/A | Department of Defense laboratories must evaluate equivalency using a process approved by the Center for Laboratory Medicine Services. | The requirement for evaluating the equivalency of qualifications for embryology laboratory directors trained outside the US has been revised to clarify acceptable records for equivalency and to add a specific clause for Department of Defense laboratories.
RLM.10250 | physical | N/A | The qualifications for embryology laboratory personnel performing ART procedures have been revised to remove "physical" from the list of qualifying degrees, and added medical laboratory science, or medical technology.
COM.01200 | non-test activities | other related information, such as methods used and types of services offered | The activity menu must be accurate for other related information, such as methods used and types of services offered, not just test activities.
| reported to | listed on | Laboratories subject to CLIA must list patient-specific results reported from the laboratory on the CAP Activity Menu.
COM.01400 | N/A | Cytogenetics and Transfusion Medicine | The qualifications of a technical supervisor or section director for Cytogenetics and Transfusion Medicine have been added.
COM.01520 | This is not permitted for ISH slides. The laboratory must not participate in formal (external) PT if the hybridization and interpretation are performed at different laboratories. Participation in formal PT would constitute PT referral. | PD-L1.DNA MMR | Laboratories that perform hybridization and ISH interpretation at different locations are now explicitly prohibited from participating in formal PT programs.
COM.01700 | It is recommended that. The CAP recommends, but does not require | The CAP recommends, but does not require | The guidance on investigating acceptable results that show significant bias or trends has been revised to state that the CAP recommends but does not require investigation.
COM.04250 | detection of DNA versus a biochemical characteristic | MALDI-TOF versus phenotypic/biochemical microbial identification | For Microbiology testing, the requirement applies when two instruments are used to detect the same analyte, excluding cases where different methods are used to identify an organism by detecting different analyte characteristics, such as MALDI-TOF versus phenotypic/biochemical microbial identification.
GEN.41096 | N/A | N/A | The 2023 requirement COM 22950 was merged into the new 2024 requirement GEN 41096.
DRA.10100 | Have at least one year of laboratory training during medical residency/fellowship; or | Have at least 20 CE credit hours in laboratory practice that cover director responsibilities as defined in the DRA checklist. | The requirement for the high-complexity testing director to have 20 CE credit hours has been added. CLIA AND AAB ROLLED THIS BACK IN JULY of 2025.
| At least one year of experience supervising nonwaived laboratory testing; or | CurrentHave current certification by a board approved by HHS***, and i. Have at least one year of experience directing or supervising nonwaived testing. | The qualification for a doctoral degree holder to be a director of a moderate complexity laboratory has been changed from supervising nonwaived testing for one year to having board certification and at least one year of experience directing or supervising nonwaived testing.
DRA.10150 | N/A | Pathologists who qualified to provide these services prior to December 28, 2024, may continue to provide these services if they have done so continuously in a CLIA-certified laboratory. | An exception for pathologists who qualified for anatomic pathology services before December 28, 2024, has been added.
DRA.10432 | N/A | N/A | This is a new requirement for directors of laboratories subject to US regulations to conduct on-site visits at least every six months.
DRA.10433 | N/A | N/A | This is a new requirement for directors of laboratories not subject to US regulations to conduct on-site visits at least once per year.
DRA.10435 | N/A | N/A | This revised requirement now directs laboratories to define a policy or agreement for the frequency of on-site visits if the director's activities are routinely conducted remotely.
GEN.20180 | N/A | in accordance with the manufacturer's instruction for use | Laboratories are now required to maintain and clean instruments and equipment in accordance with the manufacturer's instructions.
GEN.20450 | N/A | If a test system or instrument is not in use for an extended period, the laboratory defines its policy regarding preventive maintenance to be performed prior to returning the system to service. | A new requirement has been added for laboratories to have a policy for preventive maintenance on instruments not used for an extended period.
GEN.20665 | and are inspected at least every 6 months to ensure that the refrigerator and freezer units are clean and in good repair. | N/A | The requirement for inspecting refrigerators and freezers every 6 months for cleanliness and good repair has been deleted.
GEN.20780 | an adequate number of | N/A | The requirement for having "an adequate number of" centrifuges has been deleted.
GEN.20925 | a device that is used to automatically dilute or dispense liquids. | N/A | The definition of a mechanical pipetting device has been removed from this requirement.
GEN.20950 | N/A | N/A | The requirement now explicitly states that pipette verification of volume and function is an exception to the general preventive maintenance requirement.
GEN.22018 | N/A | N/A | This is a new requirement for laboratories using computer systems to have a process to check for corrupted data files, especially in a cloud-based system.
GEN.40450 | in the manual | for use in the laboratory | The requirement for the procedure manual to include the reference intervals has been changed to state that reference intervals for use in the laboratory must be established or verified.
GEN.40460 | N/A | N/A | This is a new requirement for laboratories to verify or establish reference intervals for all tests prior to reporting patient results.
GEN.40575 | N/A | or the laboratory’s defined range based on test methodology, reagent, and/or instrument system | The requirement for linearity has been revised to allow for the use of a laboratory-defined range based on the test system.
GEN.40600 | A minimum of three levels must be included in the analytical measurement range studies; two of which must be at or near the low and high values of the established or verified analytical measurement range. For qualitative assays, the manufacturer's analytical sensitivity and specificity claims must be verified. | N/A | The specific requirements for the number of levels in analytical measurement range studies and qualitative assay verification have been deleted.
GEN.40630 | N/A | The linearity verification studies performed for quantitative tests must include at least 3 levels, 2 of which are at or near the low and high values of the established or verified analytical measurement range. For qualitative assays, the manufacturer’s claims for analytical sensitivity and specificity must be verified. | This is a new requirement that specifies the number of levels for linearity studies and the verification of analytical sensitivity and specificity for qualitative assays.
GEN.40755 | When the manufacturer has not established performance characteristics for FDA-cleared/approved tests, the laboratory must verify performance characteristics prior to placing the test in service. | N/A | The requirement for laboratories to verify performance characteristics for FDA-cleared/approved tests has been deleted, as it is now covered by other requirements.
GEN.40800 | N/A | If the laboratory determines the test has been cleared by the FDA for the intended use...the laboratory is not required to perform further verification studies.. | The requirement now states that no further verification studies are needed if a laboratory determines a test is FDA-cleared for the intended use.
GEN.40875 | N/A | when the laboratory modifies the test methodology. | The requirement for verification studies now includes performing them when the test methodology is modified.
GEN.41160 | The laboratory must have a procedure for the calibration of a test or test system when the manufacturer’s instructions are not available or are incomplete. The procedure must be approved by the director.. | N/A | This requirement regarding calibration procedures has been deleted.
GEN.50500 | N/A | and the method for reporting proficiency testing (PT) results must be approved by the director or director's designee. | A new requirement has been added that the method for reporting proficiency testing results must be approved by the director or their designee.
GEN.50505 | N/A | Proficiency testing (PT) results must be reported in accordance with the PT program's instructions for reporting. | A new requirement has been added that PT results must be reported in accordance with the PT program's instructions.
GEN.50550 | The laboratory must use the same PT procedures, including the same reagent, equipment, and personnel, that are used for patient testing. | The laboratory must test PT samples in the same manner as patient specimens, including use of the same reagent, equipment, and personnel that are used for patient testing. | The wording of this requirement has been changed to state that PT samples must be tested in the same manner as patient specimens.
GEN.50555 | N/A | The laboratory's policies for handling PT samples must define acceptable testing conditions. | A new requirement has been added that laboratory policies for handling PT samples must define acceptable testing conditions.
GEN.50650 | must be approved | must be evaluated for accuracy and approved. | The requirement for PT score reports has been changed to require that they be evaluated for accuracy and approved.
GEN.50850 | A plan of corrective action must be approved by the director. | A plan of corrective action must be documented and approved by the laboratory director. | The requirement for a corrective action plan has been changed to state that it must be documented and approved by the laboratory director.
GEN.50875 | N/A | A plan of corrective action must address all PT results that were not perfect.. | A new requirement has been added that the corrective action plan must address all proficiency testing results that were not perfect.
GEN.51500 | A written plan for the ongoing review of external PT results must be approved by the director. | A written plan for the ongoing review of external PT results, including the process for evaluating and documenting the PT results, must be approved by the laboratory director. | The requirement for an external PT review plan has been revised to include the process for evaluating and documenting the PT results.
GEN.51515 | at least every 6 months | N/A | The requirement for external PT reviews to be performed at least every 6 months has been deleted.
GEN.51550 | The laboratory must investigate and document all acceptable results that show significant bias or trends. | The laboratory must investigate and document all acceptable results that show significant bias or trends. The investigation is not required when the laboratory has already identified the source of the bias or trend.. | The requirement for investigating acceptable results that show bias or trends has been revised to state that the investigation is not required if the source of the bias has already been identified.
GEN.51600 | The laboratory must have a procedure for the review of all PT samples before they are submitted. | The laboratory must have a procedure for the review of all PT samples, including the process for evaluating and documenting the PT results, before they are submitted. | The requirement for a PT sample review procedure has been revised to include the process for evaluating and documenting the PT results before submission.
GEN.51625 | N/A | N/A | This is a new requirement for the laboratory to have a procedure for the review of all PT samples before they are submitted.
GEN.51650 | N/A | N/A | This new requirement states that if a PT sample is submitted from an instrument or method not used for patient testing, this must be documented in the PT records.
GEN.51700 | All PT samples are returned to the PT provider unless the PT provider specifies otherwise. | All PT samples that are not returned to the PT provider must be stored in a manner that prevents their use for patient testing. | The requirement for returning PT samples has been changed to state that if they are not returned, they must be stored to prevent their use for patient testing.
GEN.51800 | PT specimens must be tested by the same personnel, using the same procedures, and on the same instrument as patient specimens. | PT specimens must be tested by the same personnel, using the same procedures, and on the same instrument as patient specimens. This is a new requirement for laboratories subject to US regulations. | The requirement for testing PT specimens has been clarified as a new requirement for laboratories subject to US regulations.
GEN.52200 | For laboratories with multiple sites or departments, the laboratory director must establish a policy to ensure that all PT specimens are tested at the site and department that is subject to the inspection.. | For laboratories with multiple sites or departments, the laboratory director must establish a policy to ensure that all PT specimens are tested at the site and department that is subject to the inspection. This is a new requirement for laboratories subject to US regulations. | The requirement for PT specimens in multi-site laboratories has been clarified as a new requirement for laboratories subject to US regulations.
GEN.52700 | All PT specimens are retained in a secured, locked area until they are shipped to the PT provider. | N/A | The requirement for retaining PT specimens in a secured, locked area has been deleted.
GEN.53600 | N/A | in states where state licensure is required | The requirement for a technical supervisor to have a current state license has been clarified to apply only in states where it is required.
GEN.53625 | N/A | a technical consultant | The reference to a technical consultant has been updated to reflect the new designation.
GEN.53650 | N/A | N/A | This requirement regarding the qualifications of a general supervisor has been deleted.
GEN.54200 | N/A | An individual's license must be available to the laboratory at all times. A certified copy or a copy of the license must be available on the premises. | This is a new requirement for laboratories to have a copy of a licensed individual's license available on the premises.
GEN.54750 | N/A | N/A | The qualifications for testing personnel have been deleted.
GEN.60250 | a written safety plan that is annually reviewed | a written safety plan that is reviewed annually | The wording for the safety plan review has been changed to be more concise.
GEN.60500 | N/A | N/A | The requirement for a chemical hygiene plan has been deleted.
GEN.60750 | and for personnel who may reasonably be expected to be exposed to bloodborne pathogens | N/A | The phrase "for personnel who may reasonably be expected to be exposed to bloodborne pathogens" has been deleted.
GEN.61000 | N/A | For laboratories subject to US regulations.... | The requirement for a biological safety program has been updated to explicitly apply to laboratories subject to US regulations.
GEN.70010 | Emergency eyewash and shower equipment are inspected and maintained as specified in ANSI Z358.1-2014 or the most recent revision. For plumbed systems, this includes a weekly activation check to verify operation and ensure flushing fluid is available. For self-contained units, a visual examination is performed weekly to determine if flushing fluid needs to be supplemented or changed. Plumbed systems are protected from unauthorized shutoff. Hands-free flow once activated. The equipment is capable of delivering 1.5 L per minute for 15 minutes. Flow is provided to both eyes simultaneously. Nozzles or covers to protect from airborne contaminants. | N/A | The previous requirement for emergency eyewash and shower equipment has been deleted and replaced with a new requirement.
GEN.70020 | N/A | Emergency eyewash and shower equipment must be accessible and be inspected and maintained as specified in ANSI Z358.1-2014 or the most recent revision. | A new requirement has been added for emergency eyewash and shower equipment to be accessible, inspected, and maintained according to ANSI standards.
GEN.70030 | N/A | N/A | This is a new requirement for emergency eyewash and shower equipment to be activated weekly for plumbed systems or visually examined weekly for self-contained units.